## Slide 1
Alan Lynch
Head of Wheeled Mobility and
Seating Centre Blackpool
Chairman of International (ISO), European (CEN), British (BSI) Standards Groups
for Wheelchairs and Seating
30 JUNE 2003
GSRG (Ad-hoc) Group
Informal document No. 9
(85th GRSG, 21-24 October 2003 agenda item 3.)

## Slide 2
On 1 April 2003
MDA joined with
MCA (Medicines Controls Agency)
to become the MHRA
(Medicines and Healthcare products Regulatory Agency)
MHRA is an Executive Agency of the Department of Health

## Slide 3
The MHRA (Devices) is the primary source of expertise within the UK Department of Health for matters relating to the safety, quality and performance of all medical devices.
Also on behalf of the UK Secretary of State for Health MHRA (Devices) is the UK Competent Authority for the European Medical Devices Directive (CE marking).

## Slide 4
Major areas of MHRA (Devices) workload:-
Adverse Incident Investigations
Advice & Guidance
European Affairs
Standards Activity
Product Evaluation

## Slide 5
There are over 1,000,000 wheelchair users in the UK and thousands of vehicle journeys are completed every day with few reported problems.
However, a small number of injuries and fatalities have been reported.

## Slide 6
MHRA has attempted to raise awareness and understanding for all concerned by producing two guidance documents on the Safe Transportation of Wheelchairs.
Guidance on the Safe Transportation of Wheelchairs MDA DB 2001(03) June 2001
Guidance on the Safe Use of Wheelchairs and Vehicle-mounted Passenger Lifts MDA DB 2003(03) March 2003
The documents are available as a download from MHRA website www.mhra.gov.uk and in paper form from MHRA.

## Slide 7
The UK Medical Device Regulations 2002 SI No 618
The Disability Discrimination Act 1995
European Union Directives on Vehicle Construction
Manual Handling Regulations
International Standards
All have implications for wheelchair users, carers, prescribers, transport providers, vehicle and equipment manufacturers/suppliers.

## Slide 8
In general the safety of wheelchair users in vehicles should compare with other vehicle users.
When a wheelchair user is travelling in a vehicle where other seated passengers have an occupant restraint and/or headrest then an occupant restraint and/or headrest should be provided for the wheelchair user.

## Slide 9
As part of their CE marking process manufacturers of wheelchairs must:
1. Undertake a risk analysis.
2. Ensure that the device meets all the requirements of the relevant directives.
3. Make a declaration of conformity.

## Slide 10
RISKS/HAZARDS EXAMPLES (Wheelchairs)
Start with a description of the intended user
Then identify ALL intended usage
Risks/hazards both from and to :-
User, environment, construction materials, transport, maintenance/repairs, electrical safety, fire retardency, biocompatibility, hygiene, any other potential hazards

## Slide 11
MINIMISE OR REMOVE RISKS
Consider:-
Specification/design of equipment
Instructions/Warnings
Labelling
Training

## Slide 12
In some cases standards can be used by manufacturers to show how they have reduced risks and met some of the essential requirements of the Medical Device Regulations.

## Slide 13
For transportation elements of their risk management, many wheelchair and seating manufacturers look towards the available issued (or draft) standards.
Standards (or draft standards) are available for the dynamic impact testing of,
- wheelchairs
- wheelchair tie down and occupant restraint systems (WTORS)
- seating units

## Slide 14
ISO 7176 Part 19
(issued) - Wheeled Mobility Devices for Use in Motor Vehicles
ISO 10542 Parts 1 to 5
Wheelchair Tie-down and Occupant Restraint Systems (WTORS)
(issued) Part 1 General Requirements
(issued) Part 2 Four Point Strap Type Tie-downs
(draft) Part 3 Docking Tie-down Systems
(draft) Part 4 Clamping Tie-down Systems
(draft) Part 5 WTORS for Specific Wheelchairs
ISO 16840 Part 4
(draft) Wheelchair Seating for Use in Motor Vehicles

## Slide 15
ALL THESE STANDARDS INCLUDE
A forward facing impact simulation at 48km/h @ 20g.
A lap and diagonal belt restraint for the wheelchair occupant with an above shoulder ‘B’ pillar mounting or guide.
A tie-down system for the wheelchair.
Pass/fail criteria and design/labelling for the equipment itself.
NOTE: Only ISO 10542 Part 5 specifically includes children at this time.

## Slide 16
The standards concentrate on equipment pass/fail, dummy movement and the possibility of contact with the vehicle structure or other vehicle passengers.
They do not at present include any form of injury level measurement for a wheelchair user.
The wheelchair could pass whilst the dummy occupant had been subject to high impact loading as confirmed by some of the TRL and other test results in this subject.

## Slide 17
Occupant restraints provided by the WTORS manufacturer shall:
have both pelvic and upper torso belts designed to apply forces to the occupant’s skeletal regions
function independently of the wheelchair, such that the restraint belts anchor to either the vehicle or wheelchair tiedown components so that occupant- restraint loads are not transmitted through the wheelchair
have belt restraints that can be adjusted in length without the use of tools

## Slide 18
Range of required angles for pelvic belts and locations of pelvic belt anchor points

## Slide 19
If occupant restraints include structural components for the attachment of upper anchorages or guides for upper torso belts, locations for the upper anchor points shall be provided that are:
adjustable in height so they can be positioned at or above the shoulder level of the intended user(s), or that are
located at least 1 100mm above the wheelchair ground plane

## Slide 20
Clear zones for wheelchair seated occupants

## Slide 21
Example of warning label illustrating improper positioning of occupant restraint belts

## Slide 22
Illustration of proper belt fit

## Slide 23
The vast majority of wheelchairs and WTORS now being sold in the UK have been tested against these standards.
Older wheelchairs have not been tested, but reports received so far do not reveal any major problem trends in the field when used with appropriate WTORS.

## Slide 24
Lack of a “Test Certificate” for a wheelchair based on one of these standards does not mean it is totally unsafe in all situations.
Conversely
A Test Certificate does not mean it is totally safe for the wheelchair user in all situations.

## Slide 25
The wheelchair manufacturer should give information on:
suitability for use in vehicles transport
any limitations
if suitable should give information on how to transport it safely
Also the WTORS manufacturer should give information on:
compatibility
how to use safely
any limitations in use

## Slide 26
FUTURE
- The standards are still evolving.
- The inclusion of children into these standards will start soon (only 10542 Part 5 includes children at the moment).
- Rear facing in a frontal impact and forward facing in a rear impact are also under consideration. (Incorporating information obtained so far from provisional ISO Working Group tests plus the TRL research work.)
- Rear facing will probably be covered by a WTORS standard.

## Slide 27
Adverse incidents concerning wheelchair users should be reported to MHRA.
Incidents can be reported on-line or by completing paper copies of adverse incident forms obtained from our web page at:
www.mhra.gov.uk
All adverse incident reports are registered on the MHRA (Devices) database for investigation or trending.

## Slide 28
Instructions for downloading safety warnings or guidance from www.mhra.go.uk
go to Devices
go to publications/safety warnings
choose Device Bulletins to download further copies of the Device Bulletins
choose Safety Notices to download earlier Safety Warnings
choose Device Alerts for Safety Warnings after December 2002)

## Slide 29
Old MDA Web page
www.medical-devices.gov.uk
New MHRA Web page
www.mhra.gov.uk
Wheeled Mobility and Seating Centre
E-mail: bav@mhra.gsi.gov.uk
Tel no: 01253 596000
Fax No: 01253 596177