## Slide 1: Particle Measurement Programme -Analysis of Errors Paul Quincey Richard Gilham
PMP meeting
London
8th October 2007

## Slide 2: Analysis of errors
Explanation of approach used
Application to particle number concentration measurements
Application to volatile particle removal efficiency measurements

## Slide 3: Approach
Based on GUM
Guide to the Expression of Uncertainty in Measurement
BIPM, IEC, IFCC, ISO, IUPAC, IUPAP, OIML
ISO, Geneva, First Edition 1993.

## Slide 4: GUM Universal Procedure
Y = f(X1, X2, …)
Y is to be measured
Xi are influence quantities
Each influence quantity has a standard uncertainty u(xi)
Each u(xi) has a sensitivity coefficient ci =
The combined standard uncertainty uc(y) =
The expanded uncertainty U = k uc(y)
(typically k = 2 for 95% confidence)

## Slide 5: Aims of NPL PMP report
Find significant contributions to uncertainty of results
- highlight where these need to be evaluated
- highlight where these need to be constrained
Make rough estimate of full measurement uncertainties
Cannot be fully rigorous or definitive

## Slide 6: Simplified procedure used
Clarify what is being measured
Identify factors affecting the result (“influence quantities”)
Quantify the limits on these factors (“u(xi)”)
Quantify the effect of each on the result (“ci.u(xi)”)
Conclusions and rough totals (“U”)

## Slide 7: Definition: Number concentration (in cm-3) of non-volatile particles in the nominal size range 23 nm – 2.5 μm at the PCMS inlet. The air volume is to be corrected to Standard Temperature and Pressure (273.2K and 101.33 kPa)

## Slide 8: Cinlet = CPNC x F(av)particle dilution
Eg PNC calibration accuracy
PNC drift since calibration
Presence of volatile particles
Eg nonlinearity
Reproducibility
Particles in dilution air
Low size cut-off effects
Validity of average

## Slide 9: Evaluation PNC calibration criterion 10% ~ 10% dilution particle criterion 1 cm-3 ~ 3% (assuming total dilution factor 150 and lowest Cinlet of interest 5,000 cm-3) presence of volatile particles (1%) ~ ?

## Slide 10: Low size cut off – effect on typical distribution < 1%

## Slide 11: Conclusions for Cinlet The main contribution to uncertainty is likely to be the calibration of the PNC(s) High dilution factors increase the uncertainty Reproducibility and test source stability should be routinely quantified Rough total uncertainties are 15% with single PNC route, 20% with dual PNC route

## Slide 12: Definition: The proportion of tetracontane (C40) particles, of around 30 nm diameter, which having entered the VPR inlet are removed before measurement by the PNC instrument. Method 1: Evolatile removal = 1 – (F(av)particle dilution / Fvolatile particle “dilution”) Method 2: Evolatile removal = 1 – (CPNC (heater on) / CPNC(heater off))

## Slide 13: VPR efficiency – summary Absolute accuracy is greatly limited by the small concentrations expected after removal Better uncertainty could be achieved by using Method 2 and input concentrations much higher than 10,000 cm-3 In general, absolute accuracy is not required, only a demonstration that the >99% requirement is met, and most errors will not compromise this.