## Slide 1: PMP HD_ILCE Informal Meeting at DfT London 30th March 2009
Preliminary results from VE testing of Iveco Cursor 8 at Ricardo

## Slide 2: Equipment – Instruments and Configuration
Identifier | Instrument | Particles | Size range | Measurement(s)
SPCS-19 | GPMS-CVS | solid | >23nm | total PN
SPCS-20 | GPMS-MDLT | solid | >23nm | total PN
ME_3010D | ALT-CVS | solid | >23nm | total PN
ME_3022A | additional | solid | >7nm | total PN
DMS500 | additional | Solid & volatile | 5nm -1000nm | size dist, total PN
 | CVS | MDLT
Flow rate | ~56m3/min | 60litres/min
Split ratio | 1 | 1100
Volume (1800s) | 1700m3 | 1800litres
PM flow | 60litres/min | 60litres/min
Scaling factor | 1 | ~945

## Slide 3: Gaseous Emissions
Similar results from raw and dilute measurements except for HCs and to a lesser extent CO – indicates HC background in CVS

## Slide 4: Instrument Validations
All PNCs easily met zero check criterion

## Slide 5: Instrument Validations
SPCS systems easily met system zero criterion, but Matter system failed after the first few days
  Disc wear led to high background particle numbers, so later data discarded
  Matter working on new disc coatings for greater robustness; preliminary results encouraging

## Slide 6: System Backgrounds
CVS solid particle number backgrounds much higher than partial flow
DMS shows variable accumulation mode to be present
  Solid particle background probably historic soot and low volatility HCs
    Mass background also expected
  Volatile background present as nucleation mode and also associated with accumulation mode
Solid particle background

## Slide 7: PM Measurements
General trend that CVS sampled filter masses are higher than partial flow masses
The LOD can be determined as 3 x standard deviation of the blank measurement.
  Partial flow mass method LOD: 3s = 35µg (equates to 1.1mg/kWh ETC)
  Full flow mass method LOD: 3s = 108µg (equates to 8.3mg/kWh ETC)
Partial flow system 7 times more sensitive than full flow for mass measurements

## Slide 8: PN Measurements
PN emissions trends similar between CVS systems, different to partial flow
CVS PN background much higher than partial flow
Both systems indicate major transient events
  C_WHTC

## Slide 9: PN Limit of Detection
The LOD can be determined as 3 x standard deviation of the blank measurement.
Partial flow + SPCS-20
  LOD: 3 = 15 particles/cm3 in the tunnel (after dilution correction)
  equates to ~8.7x108 particles/kWh over the ETC
Full flow + SPCS-19
  LOD: 3 = 3055 particles/cm3 in the CVS (after dilution correction)
  equates to ~1.4x1011 particles/kWh over the ETC
Based upon these results, the limit of detection for the partial flow system is 160 times lower than for the full-flow system

## Slide 10: Full flow Vs Partial Flow – Correlation between Number Data
Results from cycles with highest emissions levels similar between full and partial flow
Lowest emissions cycles results only discriminated in partial flow system
CVS Background

## Slide 11: Full flow - Correlation between Different Measurement Systems
Excellent (2%) agreement between Matter system and SPCS-19, both from CVS
Dilution factors used for comparison

## Slide 12: Full flow Vs Partial Flow – Correlation between Mass Data
No correlation between mass from different measurement systems
PM from ESC cycle similar from both CVS and partial flow
ESC data similar between CVS and MDLT

## Slide 13: Partial flow Mass Vs Partial Flow Number
No obvious correlation between mass and number, even with very low background partial flow system: mass method insensitive
Poor correlation on log-log chart
No obvious relationship between PM and PN
<2mg spread in mass
400-fold range in particle number emissions

## Slide 14: Which is the ‘best’ system?
Using CoV as the discriminator…
  Background corrected mass should be avoided
  Number is generally better than mass
Partial flow PM has the lowest average CoV
CVS PN has the lowest particle number CoV
But:
  Partial flow mass method lacks sensitivity
  CVS data has high background
Repeatability not a good metric!
Confirmation of observations in the ILCE_LD

## Slide 15: Summary
Measurements from Ricardo showed high particle and HC background in the CVS, but very low background from partial flow system
  CVS >20 years old and used for many projects including DPF regeneration
    DPF calibration work may have contaminated the system leaving particle precursors that were not removed despite the PMP preconditioning
  Partial flow dilution system new, but also much more practical to clean!
PN from CVS did not correlate well with partial flow PN except for highest emissions cycles
  CVS limit of detection for PN was ~ 1011/kWh and ~8mg/kWh for mass
PM from partial flow tunnel did not correlate with PN from partial flow or PM from full flow
  Partial flow limit of detection was <109/kWh and ~1mg/kWh for mass
CoV not seen as a good metric for assessing particle number or mass systems

## Slide 16: For consideration
If Ricardo CVS is representative of working CVS systems, then a PN limit of 1011/kWh is unfeasible due to high background levels
  1012/kWh is possible
    Both partial and full flow systems could be used at this level
If partial flow only were permitted, a much lower PN limit could be applied
DPF regeneration has not been studied, and this could lead to higher backgrounds in all dilution systems