## Slide 1: The Result of Japan’s Round Robin Engine Test: Measurement of PM and PN
<number>
3/30/2009
PMP Informal Meeting
National Traffic Safety and Environmental Laboratory
Japan Automobile Research Institute

## Slide 2: Contents
<number>
Outline and Objective
Round Robin Test Results
Understanding the Factors Affecting the Measurement Results
Summary

## Slide 3: Outline and Objective
<number>
Japan’s round robin test, the measurement of PN and PM using HD engines, was conducted at two laboratories (NTSEL and JARI).
The round robin test in Japan had the following two objectives:
  To perform testing according to Validation Exercise at NTSEL and JARI and supply data to PMP;
  To study the optimum measurement method as the PN-measurement method and filter-weight method and provide data that will help determine the measurement procedures.
Notes: <number>

## Slide 4: Specifications of Test Engine and Fuel
<number>
Engine Model | HINO J08E-TP
Configuration | Inline 6cyl, w/I.C.,T.C.
Bore X Stroke | 112 X 130
Displacement | 7.684 L
Compression ratio | 18.0
Injection System | Common Rail (Max 1600bar)
Emission Reduction Device | DPF w/Cat., Cooled EGR
Performance | 177kW / 2700rpm
Emission | 2005JP, nearly EURO V
Test Engine
Test Fuel
In conformity with RF-06-03
Notes: <number>

## Slide 5: Specifications of Full and Partial Tunnels
<number>
 |  | NTSEL (Lab. A) | JARI (Lab. B)
Full Tunnel | 1st Tunnel Diameter | 457 mm | 605 mm
 | CVS Flow | 50 m3/min | 60 m3/min
 | 2nd Tunnel Diameter | 83 mm | 83 mm
 | PM Sample Flow Rate | 67 L/min | 68 L/min
 | Dilution Air Flow Rate | 50 L/min | 15 L/min
Partial Tunnel | Tunnel Diameter | 29.4 mm | 29.4 mm
 | PM Sample Flow Rate | 67 L/min | 68 L/min
 | Split Ratio of Exhaust Flow | 1/700 | 1/1000
PM Sample | Filter Material | TX47Φ | TX47Φ
 | Temperature | 475C | 475C
Notes: <number>

## Slide 6: Particle Number Counting System
<number>
2 Horiba systems and 1 Matter system were used. One of the Horiba systems was used by both laboratories as common system.
System | Maker | Model No. | PND1 dilution ratio setting | PND2 dilution ratio setting | Tested by
SPCS1 | Horiba | MEXA-1000SPCS | 10 | 15 | NTSEL (Full) JARI (Full/Partial)
SPCS2 | Horiba | MEXA-1000SPCS | 20 | 15 | NTSEL (Partial)
Matter | Matter | MD19+ASET15 | 22 | 10 | JARI (Full/Partial)
NTSEL
Conducted simultaneous measurement with full and partial tunnels in the same test using 2 SPCSs.
JARI
Conducted measurement at the same sampling point using 1 SPCS and 1 Matter. Measurement with full tunnel and that with partial tunnel were performed in separate tests.
Notes: <number>

## Slide 7: PNCS Data to Calculate Emission
No problem with PNC provisions (cut-off point, correlation with reference meter).
As to PCRF, all systems were within the specified range; however, Fr(50nm)/Fr(100nm) ratio was below 1 in one system. Possibly problematic.
PNC cut-off point
Linearity check against electrometer
PCRF regulation
7

## Slide 8: Daily Schedule of the Round Robin Test
<number>
 | Mon. | Tue. - Fri. | Remarks
0 | Each lab’s own test for study | IFV | 
1 |  | BG | 
2 |  | C_WHTC#1~8 | 
3 |  | 10 min soak | 
4 |  | H_WHTC#1~8 | 
5 |  | CP | 
6 |  | WHSC#1~8 | 
7 | Rated 20 min. | Rated 20 min. | 
8 | Engine stop 10 min. | Engine stop 10 min. | 
9 | Preliminary JE05 | Preliminary JE05 | 
10 | Engine stop 10 min. | Engine stop 10 min. | 
11 | Full JE05 test | Full JE05 test | 
12 |  | 15 min. | 
13 |  | BG | 
IFV: Instrument Functional Verification (Daily check of PNCS, etc.) CP: Continuous Protocol (Idle 6 min.  WHSC mode9 10 min.  Engine stop 5 min.) PC: Power Curve (Not to be recorded; To be checked at rated operation before JE05 test; rated output only) SMP: Stabilization Mode Protocol (JE05 test to be defined as SMP)
Tunnel blank (TB) measurement:
Sampling time: 30 min. Measurement timing: At start and completion of daily test

## Slide 9: Contents
<number>
Outline and Objective
Round Robin Test Results
Understanding the Factors Affecting the Measurement Results
Summary

## Slide 10: Gaseous Emission
<number>
* CO & HC: Slight differences in the result among labs in some modes; good repeatability.
* NOx & CO2: Almost the same results among labs; good repeatability.

## Slide 11: PM Emission
<number>
Without tunnel blank correction
With tunnel blank correction
Before TB correction, emission levels of Lab. A were significantly higher than the others; After TB correction, emission levels of Lab. A and Lab. B became almost the same.
PM emission: JE05 > WHTC-Hot  WHTC-Cold  WHSC
Full tunnel
Partial tunnel
Without tunnel blank correction
Full tunnel
With tunnel blank correction
Partial tunnel

## Slide 12: PN (Number) Emission -Full Tunnel-
<number>
Full tunnel:
* The emission levels in WHTC-C/H of Lab. A and Lab. B were almost the same.
* The emission level in WHSC of Lab. A and that in JE05 of Lab. B were about 3 times higher, respectively.
Also, even when the same PNCS was used, the result differed between labs.
Inter-Lab test
Round-Robin test

## Slide 13: PN (Number) Emission -Partial Tunnel-
<number>
Partial tunnel:
* The emission levels of Lab. A and Lab. B were the same as the full tunnel, except for JE05.
* The emission level in JE05 of Lab. B was about double the level in the full tunnel.
Also, even when the same-model PNCS was used, the result differed between labs.
PN emission: JE05 > WHTC-Cold > WHTC-Hot  WHSC
The emission-level order of WHTC-Cold differed between PM and PN.
Round-Robin test
Same model

## Slide 14: Comparison of COV for PM
<number>
TB correction on PM lowered emission levels, causing COV to be higher.
However, TB correction reduced the standard deviation (g/kWh), resulting in smaller variations.
Without TB correction
With TB correction
Without TB correction
With TB correction

## Slide 15: Comparison of COV for PN
<number>
COV for PN: * No large difference between full tunnel and micro tunnel dilutions; around the same level. * COV was 20%-70%; almost the same level as PM without TB correction.
Full tunnel
Full tunnel
Partial tunnel
Partial tunnel
RR test
IL test
Another type PNCS
Same type PNCS

## Slide 16: PM Emission in ETC and ESC Modes
No unique tendency in the emission level and COV was observed in either ETC or ESC.
Emission
COV
16

## Slide 17: PN Emission in ETC and ESC Modes
Correlation of PN and PM was good in both ETC and ESC, similarly to WHTC Hot, etc.
COV were relatively high in ETC and ESC, which is presumed to have been caused by the conditions being not uniform due to the simple preconditioning procedures.
Emission
COV
17

## Slide 18: Effect of Tunnel Blank - Measured values -
<number>
Amount of PM trapped (mg/30min)
Average PN (#/cc/30min)
PM: TB levels of PM were high with test of the partial tunnel of Lab. A.
PN: TB levels were higher with SPCS system than Matter system in this research. TB levels of the same-model system were around the same level in any laboratory.  Presumably due to the noise from CPC and differences in the dilution method.

## Slide 19: Effect of Tunnel Blank - Converted to mode emission -
<number>
TB levels of PM were high with the partial tunnel of Lab. A, and its contribution to the emission was also high.
PM calculated from TB-PM weight
PM emission
Full tunnel
Partial tunnel
Full tunnel
Partial tunnel
PM calculated from TB-PM weight
PM emission

## Slide 20: Effect of Tunnel Blank - Converted to mode emission -
<number>
Since the effect of TB is affected by Vmix and power output of engine, the level of effect differs among modes.
* With Matter, TB contribution was about 1/100 to 1/1000; low effect on emissions.
* With SPCS, TB contribution was about 1/10 to 1/100; effect on some modes.
* The difference in the level of effect between Matter and SPCS is possibly due to the difference in the dilution method and the CPC noise level.
PN
Emission calculated from TB conc.
Emission

## Slide 21: Impact of PN Tolerance on PN Emission
<number>
Since the set dilution ratio was higher in Matter than SPCS, the level of effect was also higher in Matter.
Since the equipment specifications differed for each lab, the level of effect also differed for each test condition.
When the number of particles within the limit, 0.5 particle/cm3, was kept being measured, the calculated levels were the same as or exceeded the measured PN emission levels in some modes (WHTC-Hot, WHSC).
Regarding check procedure of R83 on pre-measurement (0.5 particle/cm3 or less when equipped with HEPA filter), the level of effect of this being measured as TB was calculated using the measured Vmix and power output.
Emission calculated from Tolerance
Emission

## Slide 22: Contents
<number>
Outline and Objective
Round Robin Test Results
Understanding the Factors Affecting the Measurement Results
  Correlation of Two Counting Systems
  Sensitiveness of the Methods: PM & PN
  Effect of Filter Material on PM Emission
Summary

## Slide 23: Correlation of SPCS-1 and Matter
<number>
Correlation of the two PNCSs was very good (coefficient of determination R2 > 0.99).
However, the linear approximation slopes had the value of about 0.8, which means the emissions measured by SPCS-1 were lower than those by Matter by about 20%.
Full tunnel
Partial tunnel

## Slide 24: Possible Factors for the Differences in the Results
<number>
Differences in the correction method between two systems
  Matter (JARI’s method)
    Particle type: NaCl particles produced by vaporization and condensation method
    Concentration: About 800010000 particles/cm3
  SPCS (Horiba’s method)
    Particle type: NaCl particles produced with atomizer
    Concentration: Varies depending on the particle size
Other factors
  Dilution method
  CPC measurement sensitivity
  …..
Fr calibration was performed on both systems according to the R83 regulations; however, there was a difference of 20% or more in their measurement results. Hence, the above factors need to be studied/discussed.
Since the possibility of the differences in the calibration method affecting the results cannot be denied, more detailed calibration procedures should be specified.

## Slide 25: Sensitiveness of the Methods
<number>
* PM: Mode ratio (Hot/Cold) of the PM weight was constant, with or without soot loading.
* PN: Ratio (Hot/Cold) of the PN count was decreased by repetition, with or without soot loading.
* It is possible that DPF and other conditions cause experimental errors.
To improve the repeatability of PN measurement, more strict test procedures are necessary.
w/ soot loading
w/o soot loading
PM
w/ soot loading
w/o soot loading
PN
The results of the WHTC Hot tests repeated after WHTC Cold tests on that day (shown against the corresponding WHTC Cold test results)

## Slide 26: Effect of Filter Material on PM Emission
<number>
* Since the effect of gas adsorption is small on Teflo filters, even the non-TB-corrected PM emissions measured with Teflo were lower than the TB-corrected emissions measured with TX40.
* Change of filter from TX to Teflo lowered PM emission levels, causing COV to be higher.
* However, the use of Teflo made variations (standard deviation in g/kWh) smaller, resulting in the emissions around the same level as TB-corrected TX-filter emissions (WHTC-Cold & Hot) or smaller (WHSC & JE05).
Emission
COV
Standard deviation

## Slide 27: Conclusion 1: Results of Repeated Tests
<number>
PM:
  Before TB correction, emission levels of the partial tunnel tended to be higher than the full tunnel, but TB correction reduced the difference between partial and full tunnels as well as the difference between labs.
  TB correction lowered the emission levels, causing the apparent COV to be higher. However, it made the standard deviation (variations) smaller.
PN:
  Comparison between labs: The PN emission levels in WHTC-Cold and Hot were the same between labs, but there were about 2- to 7-fold differences in WHSC and JE05.
  Comparison between full and partial tunnels: In Lab. A, the emission levels were the same between full and partial tunnels, but in Lab. B, the partial tunnel emissions were double the full tunnel emissions in JE05. It is important to investigate the factors affecting the results, such as the differences in the dilution ratio, test environment, etc. among labs.
COV:
  The PM COV and PN COV were mostly around the same level.
  However, for PN, since the measurement results are largely affected by engine factors such as deposition conditions of DPF, more detailed test procedures are necessary.
Correlation of PM and PN emissions:
  In the emissions from engines equipped with DPF, correlation of PM and PN emissions was small.
TB correction:
  PM: The correction reduced the difference among labs.
  PN: Depending on the equipment specs. (CVS flow rate) or test mode (power output), calculation errors that exceed the levels of measured emissions from engines are possible to generate.
  Based on the current technologies, it is necessary to give sufficient consideration to the lower limit that is measurable.

## Slide 28: Conclusion 2: Factors Affecting the Results of PN Measurement
<number>
While the correlation of the two PNCSs used in this test was good, a difference of about 20% in the sensitivity was observed between SPCS and Matter.
The difference in the sensitivity of PNC can be corrected using the sufficient correction scheme that is currently available. However, the VPR PCRF correction method varies among manufacturers.
To minimize differences in the measurement results, it is necessary to prescribe more detailed provisions on those items that would affect the measurement results, such as the correction method.
PN emissions in the repeated tests decreased continuously. Therefore, it is necessary to establish detailed test procedures for PN measurement by giving sufficient consideration to the loading conditions of DPF.

## Slide 29: Tasks for Particle Number Measurement
<number>
To measure PN with good reproducibility, it is necessary to establish strict test procedures including the preconditioning conditions, etc.
TB correction on PM is an effective method for mitigating differences among labs and/or systems.
Depending on the equipment differences or test mode, it is possible that tunnel blank/system blank-level particle concentrations will contribute to the PN emissions.
Sufficient consideration should be given to the lower limit for measurement.
In drafting the provisions on the PN calibration method, it is desirable to make modifications to the details:
  PCRF correction and its allowable range
  Strict regulations on particle species and concentration, etc. for calibration particles

## Slide 30: Appendix
<number>
Test System Diagram
Correlation of PN and PN emission
Effect of Dilution Ratio in PNCS
Effect of Introducing-Hose Temperature
Effect of Introducing-Pipe Length
Effect of Micro Tunnel Sampling Position

## Slide 31
<number>
Test System Overview (NTSEL)

## Slide 32
<number>
CFV
Dilution Air
C
HEPA
Full dilution tunnel
Test engine
Engine dynamometer
PNCS1
Matter System
ASET15
Particle Mass measurement
2nd tunnel
PNCS2
Horiba System
SPCS
Test System Overview (JARI Full Tunnel)

## Slide 33
<number>
Test System Overview (JARI Micro Tunnel)
Test engine
Engine dynamometer
PNCS1
Matter System
ASET15
PNCS2
Horiba System
SPCS
Dilution Air
Particle Mass measurement (TX)
Vent
Partial Tunnel

## Slide 34: Correlation of PM and PN emission
<number>
Correlation of PM and PN was low at the emission levels of DPF-equipped engines.
Relationship between PM and PN emissions was likely to differ among modes.
Full tunnel –Lab. A SPCS1
Full tunnel –Lab. B Matter
Full tunnel –Lab. B SPCS1
Partial tunnel –Lab. A SPCS2
Partial tunnel –Lab. B Matter
Partial tunnel –Lab. B SPCS1

## Slide 35: Effect of Dilution Ratio (PND1, PND2)
<number>
PND1
In PND1, it was indicated that changes in the dilution ratio might affect the measurement results slightly.
In PND2, the effect on the measurement results was small in the dilution ratio range under PMP Recommend.
PND2
PMP Recommend
PMP Recommend
PMP Recommend
PMP Recommend

## Slide 36: Effect of Introducing-Hose Temperature
<number>
Comments:
* In WHTC-Cold mode, the PN (full flow/partial flow) decreased when the sample-introducing-hose temperature was 50C or below.
* No effect of temperature was observed in the repeated tests in WHTC-Hot mode.
Conclusion: If the hot hose is used, there will be no effect on the particle loss.
Effect of introducing-hose heating on the particle loss

## Slide 37: Effect of Introducing-Pipe Length
<number>
Comments:
* In WHTC-Cold mode, the PN decreased by 10% compared to the Hot mode.
* In WHTC-Hot mode, no effect of exhaust pipe length was observed.
* When sample was taken from the 2nd tunnel, the PN (full flow/partial flow) increased (cause unknown).
Conclusion: No effect of exhaust pipe length was confirmed.
Differences in the particle loss depending on the introducing-pipe length and sampling position

## Slide 38: Effect of PT Sampling Position
<number>
PM (Mass) emission
PN (Number) emission
* The average of the measurement results indicated that sampling position changes tended to decrease particle emissions for both PM and PN.
* However, there were relatively large variations in the results; hence, statistically speaking, it can be judged that there is little effect of the exhaust pipe length on the PM and PN emission levels.
Mode: WHTC-Hot
Mode: WHTC-Hot

## Slide 39: Full/Partial Ratio and PM & PN Emissions
* Full/Partial COV was improved for PN; therefore, the variations observed in the PN measurements are presumed to have been caused by variations in the engine exhaust.
39