OICA PROPOSAL FOR AMENDMENTS TO REGULATION No. 83

This proposal, submitted by OICA, is intended to address issues identified after further consideration of the OICA proposal, document ECE/TRANS/WP.29/GRPE/2009/8, to amend the requirements in the Regulation No. 83 to clarify and update the specifications for the Flame Ionisation Detector (FID) hydrocarbon analyser which is used during testing of the vehicle.

At its 58th session in June 2009, GRPE agreed to defer consideration of ECE/TRANS/WP.29/GRPE/2009/8 to its next session in January 2010. However, the independent assessment of the paper that had been requested was not available for this session and the paper was therefore withdrawn by the expert from OICA.

This independent assessment has now been provided in German by ATE EL (an EU technical service based in Luxembourg) and is attached as appendix 1 to this paper.

An unofficial translation of this document has been provided by OICA and is attached as appendix 2.

The text of the document ECE/TRANS/WP.29/GRPE/2009/8 remains unchanged.

- - - - -

Analysis of the comparability of the measurement procedures for establishing HC emissions from diesel engines using heated and unheated FID

1.) Status:

According to the current status of ECE Regulation No. 83 (Supplement 8 to the 05 series of amendments entering into force on 22.07.2009), HC measurements should be conducted for compression engines with a heated Flame Ionisation Detector (FID) and for spark ignition engines with a non-heated FID.

The different measurement methods have the following background:
A small proportion of volatile components of unburnt diesel fuel can condense either before or within the measurement system and would therefore be excluded.

The continuing reduction of emissions limits along with technical measures lead to the situation that very few of these components remain therefore rendering condensation almost impossible.  On account of this, the question arises whether the necessity to use heated FID for diesel still exists.

In a proposal to amend ECE Regulation No. 83 (ECE/TRANS/WP.29/GRPE/2009/8 of 31.10.2008), the magnitude of differences between the two systems was questioned and the usage of a single system (non-heated FID for spark as well as compression ignition engines) was requested.

2.) UData (VW/Audi-2009):

In 45 measurements of different diesel engines (see annex), the HC emissions were established using both heated and non-heated FID.  The fuel used was the reference fuel described in Regulation (EC) No. 692/2008

2.1) Manufacturer and type of FIDs used in the test programme:

Audi AG Neckarsulm:

|  | Manufacturer | Type |
| --- | --- | --- |
| FID | AVL | FID 4000 LCD |
| H-FID | AVL | FID 4000 LH |

Volkswagen AG Wolfsburg:

|  | Manufacturer | Type |
| --- | --- | --- |
| FID | ABB | Multi-FID 14 |
| H-FID | ABB | Multi-FID 14 H |

3.) Analysis of the data:

3.1) “Visual” Judgement

Judged visually, the deviation of the measurement values appears small and the distribution not uneven.  The highest deviation within a pair of measurements was 2.7 % (corresponds to 0.7 mg/km) and the average deviation of the entire test programme was 0.11 mg/km.

3.2) Statistical Judgement

As a further assessment whether the test procedures can be taken as equivalent, there follows a comparison of pairs with t-test (double mean comparison for dependant samples) and the hypothesis H0 : Δ = 0.  The differences comply with a normal distribution.

3.2.1) Compilation of n = 45 deviations : di = xAi - xBi

3.2.2) Average deviation : d = -0.1124

3.2.3) Standard distribution of the deviations : sd = 0.3318

3.2.4) sd=sdn=−0.310142=0.0495

3.2.5) tprüf=dsd=0.11240.0495=2.2723

3.2.6) Determination of tcrit

|  | tcrit |
| --- | --- |
| 0.05 | 2.015 |

As ttest ≥ tcrit, the hypothesis H0 : Δ = 0 can be discounted. On the basis of a significance level of 5% this presents a significant contradiction against the acceptance of comparability of the measurement methods.

4.) Appraisal of Findings

Despite apparently similar measured vales, t-Test presents a mathematically significant contradiction against the acceptance of comparability of the measurement methods. A visual assessment of the values in the test series considers only the distribution of the differences. The concept t-Test considers the distribution of the average differences and the result of t-Tests is statistically unambiguous.

Not everything that is statistically significant is also technically relevant. For a measurement range of 0 mg/km (zero HC emissions) to theoretically 230 mg/km (HC+NOx limit for EU5 category M1 diesel vehicles), a significant average difference of 0.11 mg/km is technically irrelevant.

The proposal to amend Regulation No. 83 (ECE/TRANS/WP.29/GRPE/2009/8 of 31.10.2008) demonstrated via 400 tests that the influence of heating the FID has today no significant effect on the emissions results. This formulation leaves open whether there is no measurement technology based mathematical effect on the emissions results (contrary to the analysis of the series of measurements) or, whether a very low (but mathematically significant) average difference is not considered relevant.

Conclusion:U

The equivalence of the measurement processes cannot be taken from the statistical assessment. The average deviation as well as the individual deviation is however so small that the request to reduce the need to invest in two measurement technologies that deliver almost identical results and consolidate on one option is valid.

The usage of FID in place of H-FID delivers equivalent results.

Both processes can therefore be considered equivalent.

Attachment: Summary of Measurements

Comparison of THC Measurements from Diesel Vehicles

| Test number | Manufacturer / Site | HC-Mass emission [mg/km] |  | Absolute deviation | Deviation [%] |
| --- | --- | --- | --- | --- | --- |
|  |  | heated FID | non-heated FID |  |  |
| 200901270902 | VW/ WOLFSBURG | 24.1 | 23.4 | -0.7 | -2.7% |
| 200901270903 | VW/ WOLFSBURG | 23.1 | 22.5 | -0.6 | -2.4% |
| 200901270904 | VW/ WOLFSBURG | 20.6 | 20.2 | -0.4 | -2.1% |
| 200901280901 | VW/ WOLFSBURG | 53.6 | 53.5 | -0.1 | -0.2% |
| 200901280902 | VW/ WOLFSBURG | 36.3 | 36.3 | 0.0 | 0.1% |
| 200901280903 | VW/ WOLFSBURG | 37.5 | 37.1 | -0.4 | -1.1% |
| 200901280905 | VW/ WOLFSBURG | 22.7 | 23.0 | 0.3 | 1.1% |
| 200901280906 | VW/ WOLFSBURG | 26.6 | 26.9 | 0.3 | 1.0% |
| 200901280907 | VW/ WOLFSBURG | 34.4 | 34.0 | -0.4 | -1.3% |
| 200901290901 | VW/ WOLFSBURG | 26.3 | 26.1 | -0.2 | -0.7% |
| 200901290902 | VW/ WOLFSBURG | 20.6 | 20.6 | 0.1 | 0.3% |
| 200901290906 | VW/ WOLFSBURG | 26.6 | 26.0 | -0.7 | -2.4% |
| 200901290907 | VW/ WOLFSBURG | 35.6 | 36.1 | 0.5 | 1.3% |
| 200902090907 | VW/ WOLFSBURG | 21.6 | 21.3 | -0.3 | -1.5% |
| 200902100903 | VW/ WOLFSBURG | 23.1 | 22.5 | -0.7 | -2.9% |
| 200902100908 | VW/ WOLFSBURG | 29.7 | 30.0 | 0.3 | 1.0% |
| 200902110902 | VW/ WOLFSBURG | 19.5 | 19.0 | -0.4 | -2.2% |
| 200902120902 | VW/ WOLFSBURG | 50.9 | 50.9 | -0.1 | -0.1% |
| 200902240902 | VW/ WOLFSBURG | 16.6 | 16.6 | 0.1 | 0.4% |
| 200902250902 | VW/ WOLFSBURG | 17.2 | 17.4 | 0.2 | 0.9% |
| 200902250908 | VW/ WOLFSBURG | 17.4 | 17.9 | 0.5 | 2.7% |
| 200902260902 | VW/ WOLFSBURG | 19.9 | 20.0 | 0.2 | 0.9% |
| 200902260908 | VW/ WOLFSBURG | 24.1 | 24.2 | 0.1 | 0.5% |
| 200903040903 | VW/ WOLFSBURG | 20.1 | 20.0 | -0.1 | -0.3% |
| 200903050904 | VW/ WOLFSBURG | 27.0 | 26.7 | -0.3 | -1.2% |
| 200902260708 | VW/ WOLFSBURG | 20.7 | 20.5 | -0.2 | -1.1% |
| 200903020707 | VW/ WOLFSBURG | 22.5 | 23.0 | 0.4 | 2.0% |
| 200903020708 | VW/ WOLFSBURG | 24.7 | 24.4 | -0.2 | -0.9% |
| 200903030703 | VW/ WOLFSBURG | 21.2 | 21.1 | -0.1 | -0.3% |
| 200903030704 | VW/ WOLFSBURG | 21.6 | 21.4 | -0.2 | -1.0% |
| 200903030709 | VW/ WOLFSBURG | 21.0 | 20.8 | -0.3 | -1.4% |
| 200903030710 | VW/ WOLFSBURG | 21.2 | 20.8 | -0.4 | -1.8% |
| 200903040703 | VW/ WOLFSBURG | 25.0 | 24.6 | -0.4 | -1.8% |
| 200903040704 | VW/ WOLFSBURG | 22.7 | 22.4 | -0.3 | -1.2% |
| 200903040709 | VW/ WOLFSBURG | 21.8 | 22.0 | 0.3 | 1.1% |
| 200903040710 | VW/ WOLFSBURG | 21.5 | 21.5 | 0.0 | -0.2% |
| 200903050702 | VW/ WOLFSBURG | 22.3 | 22.2 | -0.1 | -0.3% |
| 200903050703 | VW/ WOLFSBURG | 25.7 | 25.6 | -0.2 | -0.6% |
| 200903050706 | VW/ WOLFSBURG | 20.2 | 20.4 | 0.2 | 0.7% |
| 200903050707 | VW/ WOLFSBURG | 21.8 | 21.5 | -0.3 | -1.2% |
| 200903060703 | VW/ WOLFSBURG | 24.3 | 23.9 | -0.4 | -1.6% |
| 200903060704 | VW/ WOLFSBURG | 24.5 | 24.1 | -0.4 | -1.7% |
| 6-2009-0306 | Audi/ NECKARSULM | 22.2 | 22.8 | 0.6 | 2.6% |
| 6-2009-0311 | Audi/ NECKARSULM | 22.1 | 22.5 | 0.4 | 1.9% |
| 6-2009-0317 | Audi/ NECKARSULM | 21.7 | 21.2 | -0.5 | -2.3% |

- - - - -