## Slide 1: WLTP DTP PM-PN Subgroup Activities, Progress & Future Plans
5th DTP Meeting
Dubendorf
12th – 14th April 2011
Chris Parkin

## Slide 2: Outline
      Meetings
      Closed Issues
      Open Issues
      Broader Issues
      Future Plans

## Slide 3: Outline
      Meetings
      Closed Issues
      Open Issues
      Broader Issues
      Future Plans

## Slide 4: Subgroup Meetings
6th Web Conference – 31st January
  Large number of PM sampling issues closed
  Weighing group reported data received on temperature and humidity control. More data on Teflo filter weighing welcome
  Regeneration small group tasked with developing an outline for regen experiments to be taken during validation testing
7th Web Conference – 3rd March
  Number of PM sampling and weighing issues closed
  Regeneration small group presented summary of available studies on regeneration PN measurements. Tasked with developing DPF loading procedures
  PN provisional solutions presented on several issues
  OIL version 6

## Slide 5: Subgroup Meetings
8th Web Conference - 7th April
  Numerous PM sampling, weighing and PN issues closed
  JASIC comments on issues discussed
  OIL version 7
9th (Face to face) Meeting – early May
  Plan to review and update consolidated draft text

## Slide 6: Outline
      Meetings
      Closed Issues
      Open Issues
      Broader Issues
      Future Plans

## Slide 7: Closed Issues
PM Sampling
A1: CVS/Filter face temperature - <52C at filter face
A3: Avoidance of condensation in dilution/sampling systems – Lab Processes Sub-Group proposed text accepted
A4: Transfer tube coupling sealing materials – elastomers to be as thermally stable as possible and not be used to bridge exhaust to transfer tube connection.
A5: Transfer tube – Japanese spec on transfer tube insulation adopted. Heating permitted as an alternative
A6: Remote mixing tee – Permit use of RMT with conventional dilution tunnel
A7: Permissible variation in static pressure at exhaust tailpipe – Lab Processes Sub-Group proposed text accepted
A10: Dilution tunnel pressure at mixing point – Specification not required controlled by above

## Slide 8: Closed Issues
PM Sampling cont
A11: Use of heat exchanger for PM sample flow control – Lab Processes Sub-Group proposed text accepted
A12: Dilution system schematics – To be updated to (optionally) permit RMT and secondary dilution on all systems
A13: Pre-classifier –Clarify that shrouded probe or classifier are alternatives. Pre-classifier need not be immediately before filter holder
A14: Secondary dilution – Permit subject to same accuracy and proportionality requirements as for single dilution
A17: Dilution system residence time - Lab Processes Sub-Group proposed text accepted
A18: Molar based PM calculations – Issue is for Lab Processes Sub-Group

## Slide 9: Closed Issues
PM sampling cont.
A19: Verification of proportionality of sampling & P, T and flow senor requirements – Use US part 1065 requirements for sensor specification only
A21: Measurement specifications – T +/-1C, 15s rise time, Flow 2.5% of point, 1.5% of maximum
A22: Dilution air temperature – >95% of points at 25 +/-5C (measured as close as possible to mixing point if possibility of change from inlet source)
A23: Secondary dilution transfer tube length – secondary dilution air to be injected as close as possible to where transfer tube leaves primary dilution tunnel
A24: Particle sampling system bend radii – WLTP should specify radius of curvature should be “as large as possible”
A25: Sampling pump location – Need not be on filter holder

## Slide 10: Closed Issues
PM sampling cont.
A28: Dilute exhaust gas thermocouple type – Type need not be specified
A29: PM back-up filter for artefact subtraction – Not permitted
A30: PM probe location on CVS centreline – Not necessary
A31: PM probe diameter – 8mm min internal diameter as per WHDC
A32: Isokinetic sampling – Outside of WLTP Phase 1 scope
A33: PM probe to filter length - <2m
A34: Stability of temperature at flow meter – Fixed temperature not required where real time temperature correction (at 1Hz) used
A35: Recommended minimum filter loading – Recommend maximising PM loading but without specific recommended mass
A36: Filter face velocity – 20-105cm/s

## Slide 11: Closed Issues
PM sampling cont.
A39: PM filter holder design/stain area – At least 1075mm2 stain area (37mm diameter)
A40: PM sampling vacuum side leak check – Not mandatory
PM weighing
B1: Weighing room conditions – 22 +/-2C (+/-1C recommended), rh 45 +/-8%, dewpoint <10.5C
B4: Static neutralisation – recommend filter be neutralised. Recommend avoiding use of grounding straps and tweezers.
B5: Microbalance calibration – every 12 months to a traceable national standard. Zero & span check at the start of each session

## Slide 12: Closed Issues
PN
D2: VPR volatile removal efficiency specification – Retain current specification
D3: Use of second by second particle concentration reduction factor (pcrf) – retain current use of calibration pcrf applied to average particle concentrations over the test
D6: VPR pcrf validation check – permit use of polydisperse aerosol with 50nm mode
D7: VPR dilution air filtration and leak check requirements – retain Reg 83 requirements
D8: thermal treatment of VPR calibration aerosol – specify that calibration aerosol should be thermally stable at VPR operating temperatures
D9: VPR evaporation tube setpoint – 350 +/-10C, 0.25-0.4s residence
D10: PNC flow check - Retain 5% tolerance

## Slide 13: Outline
      Meetings Held
      Closed Issues
      Open Issues
      Broader Issues
      Future Plans

## Slide 14: Open Issues
PM sampling
A2: CVS temperature during DPF regeneration - <190°C subject to confirmation from validation testing
A8: Dilution air filtration efficiency – proposed 99.95%, but JASIC prefer 99.97%
A9: PM/PN tunnel air background correction – proposed to permit up to specified maxima. EPA & CARB can only accept dilution air, and not tunnel, background correction
A15: Tunnel pre-conditioning procedure – proposed to retain recommendation of 20 mins at 120kph, JASIC view is that this may be insufficient.
A16: Cooling to achieve sample temperature – do not explicitly refer to cooling in text. Awaiting draft text.
A20: Minimum data logging rates – 1Hz
A26: Number of PM filters – Propose one for cold start and for all hot start phases. Depends on cycle structure and weighting factors

## Slide 15: Open Issues
PM sampling cont
A27: Verification of proportionality of sampling – proposal is to document on system commissioning and as required by authority
A35: Location of PM filter isolating valve – proposal is for isolating valve downstream of filter. Need to review if plug-in HEV testing requires repeat tests
A38: PM filter media – permit both PTFE coated and membrane filters and note that membrane more susceptible to static charge. Need input from filter manufacturers on efficiency specification
PM weighing
B1: microbalance specification – 1g resolution, 2g precision proposed. JASIC prefer tighter specification
B2: sample filter conditioning - >1 hour conditioning pre and post-test. JASIC suggest longer minimum post test conditioning likely to be required

## Slide 16: Open Issues
PM weighing
B6: reference filter acceptance criteria – reviewing data
Regeneration
C1: whether to take PN measurement during regen - limited available studies suggest additional PN during regen is volatile material excluded by VPR and PNC cut-off. Proposal is to specify additional experimental measurements to be taken during validation to confirm applicability of PN measurement system. Current regen test procedures (repeating cycles, from hot start, until regen complete) likely to give reduced PN as reduces weighting of cold start.
Particle Number
D1: VPR minimum solid particle penetration efficiency – 70% at 100nm proposed, plus current max ratios for pcrf30nm:pcrf100nm and pcrf50nm:pcrf100nm

## Slide 17: Open Issues
Particle Number
D4: PNC calibration frequency – proposal is 13 month calibration interval, with new requirement to monitor counting efficiency versus measurements with other PNCs, reference PNCs or reference vehicles or replace PNC wick every 6 months
D5: PNC calibration aerosol – calibration factor to be applied where CAST aerosol used to check PNC cut-off. CAST particles used to check PNC slope should be >70nm diameter. JASIC prefer a single aerosol

## Slide 18: Outline
      Meetings Held
      Closed Issues
      Open Issues
      Broader Issues
      Future Plans

## Slide 19: Broader Issues
A5: Additional Pollutants sub-group to consider whether exhaust transfer tube insulation specified is adequate for water soluble additional pollutants, e.g. carbonyls, or whether heating required
A9: DTP decision needed on whether to permit tunnel background correction.
  Prohibit (EPA and CARB view)
  Permit (supported by other PM-PN subgroup members)
  Specify procedure only, leaving decision on whether to allow it for regional implementation
DPF active regenerations on test should be triggered, otherwise some method of indicating regeneration has started will be needed.
Where do we report recommendations for future research e.g. further examination of correcting PM for artefact?

## Slide 20: Issues for investigation in validation
DPF active regeneration measurements
  Is PN measurement system capable of excluding worst case volatile material during DPF regenerations
  CVS and filter face temperatures during passive regenerations?
Tunnel pre-conditioning
  Is 20 minutes at 120kph sufficient?

## Slide 21: Outline
      Meetings Held
      Closed Issues
      Open Issues
      Broader Issues
      Future Plans

## Slide 22: Future Plans
May 2011
Face to face meeting
  Review and update consolidated draft text
June 2011
Objectives
  Reach final agreement on all Open Issues
  Finalise PM-PN elements of GTR text