## Slide 1: Sub<23 nm Particles Key Messages
Document discussed in the PMP meeting 08 Jan 2014
Informal document GRPE-68-21
(68th GRPE, 7-10 January 2014, agenda item 7)
Transmitted by the PMP Chair

## Slide 2: Overview
Particles < 23nm are partially measured today
The system today works well and is applied across the industry
  Light duty
  Heavy duty
  Extension to NRMM underway

## Slide 3: Particle size / composition
Existing technologies (eg GDI) do not have significantly smaller particle size range than diesels with DPF
Some particles < 23nm on GDI sometimes identified but only at certain specific conditions. The percentage not measured is typically <60% (on average 30-40%) and comparable to diesels (<20%)
We do not know what these particles are – physically or chemically – more work required

## Slide 4: Measurement Equipment - PNC
PNC could be adapted to measure particle size down to 10nm
Current PNC could NOT be adjusted to measure below 10nm. Different model is necessary. It is not as simple as plug and play.

## Slide 5: Measurement Equipment - VPR
VPR Evaluated as not robust to measure down to either 10 or 3 nm today (volatile artifacts are common)
Alternative could be catalytic stripper
Catalytic stripper has also some issues to overcome
Calibration of VPR and determination of transfer functions / losses will be a further issue
Uncertainty of measurement expected to increase when solid sub-23nm particles exist.

## Slide 6: Calibration
ENV-02 recommended ‘not a soot like aerosol’ for calibration – but has not suggested what should be done instead.
Calibration significant challenges for reference CPC (would need to be calibrated to lower particle size). Potentially all calibration to be done by electrometer – fundamental

## Slide 7: Recommendations / Next Steps
No urgent need identified to revise measurement to < 23nm today.
New / emerging technologies to be kept under review regarding <23nm particles.
EU Com requested PMP to continue to develop a robust < 23nm method in case of future need.
  PMP will develop a plan and timeline
Guidelines to avoid artefacts to be developed when measuring <23nm particles
With the above recommendations, the work on regeneration measurements for LD vehicles can start.